## Slide 1: COSI Transmit: Open Source Soft- and Hardware Transmission System for traditional and rotating MR
Christian Blücher1, Haopeng Han1, Werner Hoffmann2, Reiner Seemann2, Frank Seifert2, Thoralf Niendorf1,3,4 and Lukas Winter1
1 Berlin Ultrahigh Field Facility (B.U.F.F.), Max-Delbrück Center for Molecular Medicine in the Helmholtz Association, Berlin, Germany
2 Physikalisch Technische Bundesanstalt (PTB), Berlin, Germany
3 Experimental and Clinical Research Center, a joint cooperation between the Charité Medical Faculty and the Max Delbrück Center for Molecular Medicine in the Helmholtz Association, Berlin, Germany
4 MRI.TOOLS GmbH, Berlin, Germany

## Slide 2
Speaker Name: Lukas Winter
I have no financial interests or relationships to disclose with regard to the subject matter of this presentation.
Declaration of Financial Interests or Relationships

## Slide 3: Motivation
Magnetic Resonance Imaging (MRI) is an essential medical diagnostic tool that is beyond the reach of many patients throughout the world1
A goal of the open source imaging initiative (OSI²) is to collaboratively build an affordable open source MR for research, education and the clinic
In this work we present our efforts towards an affordable open source transmission and reception system for traditional and rotating MR: spectrometer, RF power amplifier, T/R switch, LNA and RF coil
1 World Health Organization (WHO), (2013, 25. Feb.), Essential Health Technologies: Medical Equipment – Data by Country; Available: http://gamapserver.who.int/gho/interactive_charts/health_technologies/medical_equipment/atlas.html

## Slide 4: Towards an Open Source Desktop MR
Rotated permanent magnets can be used for spatial encoding [4-5]
Combining homogeneous main magnet with a linear gradient using permanent magnets [6-7]
Main magnet (B0=0.09T): 1000ppm -> (40x64x36)mm³
Linear Halbach gradient
measurement
simulation
[4] Cooley CZ, et al., Magn Reson Med, 2015, 73:p872-883 [5] Cooley CZ, et al., ISMRM, 2016, #3556 [6] Winter L, et al., ISMRM, 2016, #3586 [7] Blümler P, Concepts Magn Reson Part B Magn Reson Eng, 2016

## Slide 5: Spectrometer Hardware
Published spectrometers require custom circuitboard design6-8
A software defined radio (SDR) is a programmable hardware device that allows the realization of radio/communication system in software and has recently been used as an MR spectrometer9
GNU Radio compatible SDRs allow to separate software developments of MR imaging techniques from hardware developments
Utilization of a USRP110 equipped with 2 TX- and 1 RX-daughterboard
System functionality of spectrometer is defined in a GNU Radio Flowgraph utilizing the software package gr-MRI9
6 Asfour A, et al., Sensors, 2015; 13:16245-16262
7 Takeda K, J. Magn. Reson., 2008; 192:218:229
8 Chen HY, et al., J. Magn. Reson., 2015; 255:100-105
9 Hasselwander CJ, et al., J. Magn. Reson., 2016; 270:47:55
10 Ettus Research, USRP1, (2016, 18 Oct.); Available: https://www.ettus.com/product/details/USRPPKG

## Slide 6: Spectrometer
Schematic of the implemented SDR spectrometer

## Slide 7: Results – Spectrometer
RF-pulses, RF power amplifier un-/blanking signals, data recording and stepper motor control signals can be generated
Pulse length and amplitude, readout-length and –delay can be freely chosen
Arbitrary shaped RF-pulses can be generated (f=0-30MHz)
Sinc-pulse
Signal to control stepper motor
RF-pulse, unblanking and readout window

## Slide 8: Results - RF Power Amplifier
Frequency Range: 1.8-54MHz (B0=0.043-1.27MHz)
Power: 1kW
Driver1: two stage linear amp, Pout=5W, 33-34dB, 12V
Amp2: class AB, LDMOS (BLF188XR), Pmax=1400W, 50V
Filter3: 6-segment selectable low pass filter
1http://www.box73.de/product_info.php?products_id=1899&XTCsid=86316710323ff4c b4d6e8857b9a31277 2http://www.w6pql.com/1_kw_sspa_for_1_8-54_mhz.htm
3http://www.w6pql.com/a_1.5kw_lpf_for_160-6m.htm
Notes: Driver: two stage linear amplifier design, qrp, 33-34dB gain, Pout=5W, 12V dc
Amp: classAB, LDMOS (BLF188XR), 50Vdc
Filter: 6-segment, selectable low pass filter

## Slide 9: Results – RF Power Amplifier
Input power Pin plot versus output power Pout of RFPA  RFPA shows good linearity over investigated frequency range
Chart: Pin vs. Pout
| 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9 | 9.5 | 10 | 10.5 | 11 | 11.5 | 12 | 12.5
4.3MHz (B0=0.1T) | 86.3298 | 193.2578 | 317.52 | 426.32 | 551.12 | 677.12 | 768.32 | 865.28 | 968 | 1039.68
8.3MHz (B0=0.2T) | 76.1378 | 185.4738 | 288 | 392 | 512 | 591.68 | 768.32 | 832.32 | 933.12 | 1003.52
12.8MHz (B0=0.3T) | 57.6738 | 135.4658 | 208.08 | 288 | 369.92 | 462.08 | 537.92 | 619.52 | 677.12 | 768.32 | 832.32 | 898.88 | 968 | 1076.48
21.3MHz (B0=0.5T) | 47.8242 | 99.405 | 155.9378 | 216.32 | 269.12 | 327.68 | 380.88 | 438.08 | 474.32 | 551.12 | 591.68 | 633.68 | 706.88 | 737.28 | 768.32 | 832.32 | 865.28 | 898.88 | 933.12 | 968
42.6MHZ (B0=1.0T) | 30.2642 | 69.3842 | 107.4578 | 159.1328 | 204.02 | 246.42 | 283.22 | 332.82 | 386.42 | 432.18 | 480.5 | 531.38 | 584.82 | 612.5 | 669.78 | 714.42 | 752.72 | 752.72 | 784.08 | 848.72 | 848.72 | 915.92 | 950.48 | 950.48 | 1021.52
Pout=Pin*10^5 | 50 | 100 | 150 | 200 | 250 | 300 | 350 | 400 | 450 | 500 | 550 | 600 | 650 | 700 | 750 | 800 | 850 | 900 | 950 | 1000
 RFPA shows good linearity over the investigated frequency range

## Slide 10: Results – RF Power Amplifier with blanking circuit

## Slide 11: Results – RF Power Amplifier without blanking circuit
SDR output
RFPA output

## Slide 12: Results - RF Coil
Electromagnetic simulations (CST MWS 2012)
Solenoid RF coil to adapt to Halbach B0 field distribution
Adaptation of RF coil length (l) and number of turns (N) in order to reach a homogeneous B1+ field distribution within a sphere (d=70mm)
f=3.63MHz, dinner=96mm based on prototype Halbach magnet
AWG 20 copper wire as an electrical conductor
Final design: l=200mm, N=20

## Slide 13: Results – RF Coil Simulations
B1+-field inhomogeneity within the sample <3.4%
B1+ amplitude at the RF coil center of 2.7 mT/

## Slide 14: Results – RF Coil Measurements
Measured Q-factors: Qunloaded = 94  38kHz bandwidth Qloaded = 77  47kHz bandwidth
B1+-field measurements with pickup loop positioned along straight line at one end of RF coil
Pickup loop calibration in known B field of TEM cell11
RF coil measurements
TEM Cell
Pickup loop
calibration
11 Klepsch T, et al., Biomed Tech, 2012; 57:119-122

## Slide 15: Results - Comparison Measurements vs. Simulations
Absolute and normalized comparison of simulated and measured B1+ values of RF coil
Absolute difference: -5.1±4.2% (mean±std)
Normalized difference: 1.2±4.5% (mean±std)

## Slide 16: Result – COSI Transmission/Reception System
SDR:
  f=0-30MHz
RFPA:
  f=1.8-54MHz
  Pout=1kW (peak)
  RF coil:
  l=200mm
  d=96mm
  T/R switch12:
  f=3-5MHz
  Pin_max=1.3kW
  Preamplifier13:
  f=150kHz-30MHz
  Gain=18-20dB
  Noise figure=1-2dB
12 NMR Service GmbH, Erfurt, Germany 13 LNA4HF, 9A4QV, (2016, 15. Nov.) Available: http://lna4hf.blogspot.de/
Desktop MR (B0=0.08T) cost:
~160 USD for magnet
~1600 USD for 1kW RFPA
~1000 USD for SDR
~400 USD for T/R switch
~30 USD for LNA
<4000 USD for complete system

## Slide 17: Conclusion
SDR is capable of producing arbitrary shaped RF pulses, un/blanking the RFPA, recording acquired data and controlling rotated gradient fields (f=0-30MHz)
RFPA (f=1.8-54MHz, Pmax=1kW)
All project related files will be made available open source
Step by step towards a fully open source desktop MR system for research, education and the clinic

## Slide 18
Acknowledgement
Lukas.winter@ptb.de